## Slide 1
Water quality monitoring for
Legionella pneumophila and Escherichia coli

## Slide 2
Material of Currently Used Filters
0.45 um – PolyEtherSulfone (PES)
Polyethersulfone (PES) membrane for aqueous solutions provides removal of fine particles, bacteria, and fungi making it a versatile membrane for applications such as sample preparation, sterile filtration and infusion therapy.
0.2um - Regenerated Cellulose (RC)
Comparison Between 0.2um Filter and 0.45um Filter
Notes: RC - Regenerated cellulose (RC): made from pure cellulose without wetting agents. Chemical resistance to a wide variety of solvents. High wet strength. Hydrophilic, so suitable for aqueous and organic samples. Very low protein binding capacity. pH range ~3-12. PES - Hydrophilic, stable in low pH, have low levels of extractables, and exhibit low protein binding, making them suitable for many aqueous and organic solvents. PES membranes allow higher liquid flow than PTFE. Temperature resistant. pH range ~3-14 (sometimes quoted as 1-14).
https://www.emdmillipore.com/Web-US-Site/en_CA/-/USD/ShowDocument-Pronet?ProductSKU=MM_NF-GSWP04700&DocumentId=201306.9737.ProNet

## Slide 3
Comparison of Extraction protocol overview
Bacterial Culture
Pick colony (s) into 3mL LB
Grow at 37degC, 200 rpm for ~16 hr
Volume aliquote for extraction
Treatment 1 Reagent 1+Proteinase K
56degC for 1-3 hr.
Treatment 2 Reagent 2
70degC for 10 mins.
Treatment 1
Lysis buffer 1
80degC for 10 mins.
Est. ~20-30 mins.
QIAamp DNA mini kit consumables
Bacterial gDNA Extraction: QIAamp DNA mini kit (column)
QIAamp DNA mini kit consumables
Est. ~20-30 mins.
gDNA Quantification
Nanodrop
Qubit
Bacterial gDNA Extraction: Prepared Lysis Buffer + column
* 1 heating temperature for 10 mins
*2 heating temperature for =< 75mins
1
2
1
To have substantial amount of e. coli for testing.
E Lysis buffer:
NaOH
Ethanol
EDTA
Add Ethanol
3

## Slide 4
Spiked in Extraction protocol (Syringe Filter)
Bacterial Culture
Pick colony (s) into 3mL LB
Grow at 37degC, 200 rpm for ~16 hr
Volume aliquot for extraction
1. Prepare 10ml MiliQ water + Bacterial Culture
Either 0.2um or 0.45 um filter
QIAamp DNA mini kit consumables
1
2
To have substantial amount of e. coli for testing.
3
2. Filter water sample through syringe filter
3. Filter E lysis buffer through the same filter
455ul vs 1ml of Lysis Buffer
4
100ul of sample is eluted out

## Slide 5
Comparison Between 0.2um Filter and 0.45um Filter
Note: Using 0.4um Filter flushed with 1ml of lysis buffer gave the best results but further optimization should be carried out.
Threshold : 100
Notes: 20ul of culture was spiked into 4ml of culture. = 100ul culture. To compare dna yield extracted, the volume of lysis tested was 455ul and 1ml. Conclusion: Using 0.4um Filter flushed with 1ml of lysis buffer gave the best results. Further optimization can be carried out by increasing the contact time of the lysis buffer with the membrane. (Use a larger volume of lysis buffer or let the lysis buffer stay on the membrane for a certain time before flushing out)

## Slide 6
Comparison Between E Lysis Variations
80oC Has lower signal
RT Has lower signal
RT Has lower signal
Original E Lysis Buffer
QIAmp Kit
E Lysis Buffer using Conc. NaOH
Original E Lysis Buffer + SDS
Original E Lysis Buffer + SDS + NP40
E Lysis buffer:
NaOH
Ethanol
EDTA
Threshold : 100
Notes: Conclusion: Adding the different chemicals did not make significant changes to the original lysis buffer in terms of extraction efficiency. Heating the lysis buffer at 80oc shows a slightly lower ct as compared to room temp.

## Slide 7
Calculated Total Count:
= 125
Bacterial Counting: 1 µL, 10 µL, 100 µL volume of E .coli stock
Plating volume:
10µL E.coli stock + 90µL Media
Plating volume:
100µL E.coli stock
Plating volume:
1µL E.coli stock + 99µL Media
Estimated Total Count:
= 125 * 100
= 12500
49
47
63
30
36
54
31
Calculated Total Count:
= 4* (63+49+47+54+30+31+36)
= 4*310
= 1240

## Slide 8
Comparison Between E Lysis Variations
Original E Lysis Buffer
QIAmp Kit
E Lysis Buffer + Spiked sample w 0.22um filter
Experiments run prior.
Threshold : 100
E Lysis Buffer (Heated @80oC) + Spiked sample w 0.22um filter
Notes: Conclusion: Heating the lysis buffer before flushing did not make significant differences to the amount of E. coli captured.

## Slide 9
Calculated Total Count:
= 4* (7+9+19+11+8+15+8)
= 4*77
= 308
Bacterial Counting: 1 µL, 10 µL, 100 µL volume of E .coli stock
Plating volume:
10µL E.coli stock + 90µL Media
Plating volume:
100µL E.coli stock
Plating volume:
1µL E.coli stock + 99µL Media
Estimated Total Count:
= 308 * 10
= 3080
Estimated Total Count:
= 308 * 100
= 30800

## Slide 10
Comparison of Extraction protocol: qPCR detection of UidA (E. coli)
E Lysis buffer
EC 100 µL
QIAamp
E Lysis buffer
EC 10 µL
QIAamp
E Lysis buffer
EC 1 µL
QIAamp
E Lysis buffer
EC 0.1 µL
QIAamp
E Lysis buffer
EC 0.01 µL
QIAamp
E Lysis buffer
EC 0.001 µL
QIAamp
E Lysis buffer
EC 0.0001 µL
QIAamp
Note:
qPCR detection demonstrated that the extraction efficiency of E lysis protocol with reference protocol (QIAamp DNA mini kit) were comparable with ~1 ct difference.
Highest titer
Lowest titer

## Slide 11
Subsequent Plans:
Optimization of Prepared Lysis buffer
1. Temperature: RT vs 70degC (10mins) vs 80 degC (10mins) vs 90 degC (10mins)
2. Modification (Addition of detergent etc.)
Advantages of Prepared buffer protocol:
1. Flexible in modification of lysis buffer for optimization
2. Shorter duration of extraction from bacterial cells (10 mins instead of 75 mins)
3. 100 µL of sample (No need to centrifuge for pellet) can be processed directly.
4. Currently, 1 heating temperature (80 degC) as compared to reference protocol (56 & 70 degC).
Disadvantages of Prepared buffer protocol:
1. May be inefficient in extracting cell pellet with E Lysis buffer; ~50% (nanodrop)
qPCR difference
Direct Sample extraction vs Filtered Sample Extraction;
1. Syringe membrane filter (0.2um) may be used to trap Bacterial cells (I.e. E. coli) from water sample (Spiked-in samples).
2. Lysis of cells within the membrane filter using modified buffer.
3. Determine % of yield loss.
EC pellet
QIAamp
E Lysis
buffer