## Leg-Stds
Serogroup 1
16s - All Legionella
Standard HEX Ct value
Sample con. | Concentration ( Copies/ul) | Log10 Copies/ul |  | R1 | R2 | R3 | Mean value | [6 empty] | Concentration ( Copies/ul) | Mean value (ct)
1000 | 1414 | 3.1504 |  | 24.43 | 24.4 |  | 24.415 | [6 empty] | 511 | 24.415
100 | 130 | 2.1139 |  | 28.15 | 28.13 | 27.97 | 28.0833 | [6 empty] | 50.57 | 28.0833
10 | 11 | 1.0414 |  | 31.66 | 31.61 | 31.44 | 31.57 | [6 empty] | 5 | 31.57
1 | 1 | 0 |  | 34.74 | 34.39 | 35.61 | 34.9133 | [6 empty] | 0.44 | 34.9133
|  | Amplification Efficiency = 10^(-1/slope) - 1 |  |  | 99.92%
|  | linear regr.: |  | k | -3.3238
|  |  |  | d | 34.9852
y = kx + d
Serogroup 1
mipT1 - Serogroup 1-15
|  |  |  | Standard FAM Ct value | [9 empty] | Concentration (Copies/ul) | Mean value (ct)
Sample con. | Conc. (Copies/ul) | Log10 Copies/ul | Ct | R1 | R2 | R3 | Mean value | [6 empty] | 176.7 | 24.725
1000 | 1414 | 3.1504 |  | 24.8 | 24.65 |  | 24.725 | [6 empty] | 15.87 | 28.1633
100 | 130 | 2.1139 |  | 28.2 | 28.26 | 28.03 | 28.1633 | [6 empty] | 1.37 | 31.3633
10 | 11 | 1.0414 |  | 31.88 | 31.13 | 31.08 | 31.3633 | [6 empty] | 0.15 | 34.7967
1 | 1 | 0 |  | 34.91 | 34.8 | 34.68 | 34.7967
|  | Amplification Efficiency = 10^(-1/slope) - 1 |  |  | 106.5%
|  | linear regr.: |  | k | -3.1747
|  |  |  | d | 34.7669
y = kx + d
Serogroup 1
WZM - Serogroup 1
|  |  |  | Standard TexasRed Ct value | [9 empty] | Concentration (Copies/ul) | Mean value (ct)
Sample con. | Conc. (Copies/ul) | Log10 Copies/ul | Ct | R1 | R2 | R3 | Mean value | [6 empty] | 176.9 | 24.895
1000 | 1414 | 3.1504 |  | 24.93 | 24.86 |  | 24.895 | [6 empty] | 16.77 | 28.4567
100 | 130 | 2.1139 |  | 28.38 | 28.65 | 28.34 | 28.4567 | [6 empty] | 1.43 | 31.95
10 | 11 | 1.0414 |  | 31.97 | 32.01 | 31.87 | 31.95 | [6 empty] | 0.137 | 35.45
1 | 1 | 0 |  | 36.2 | 34.67 | 35.48 | 35.45
|  | Amplification Efficiency = 10^(-1/slope) - 1 |  |  | 99.23%
|  | linear regr.: |  | k | -3.3406
|  |  |  | d | 35.4542
y = kx + d
Unknown Samples Concentration
|  |  | Sample | Ct value HEX | Ct value FAM | Ct value TxRed | Log10(conc.) HEX | Concentration HEX (cop/ul) | Log10(conc.) FAM | Concentration FAM (cop/ul) | Log10(conc.) TxRed | Concentration TxRed (cop/ul)
|  | 1 | Water Tank | 28.44 | 39.11 | 31.08 | 1.9692 | 93.1528 | -1.368 | 0.042852 | 1.3094 | 20.3894
|  | 2 | Passed Tap | 24.745 | 29.74 | 32.89 | 3.0809 | 1204.7022 | 1.5834 | 38.318 | 0.76759 | 5.8558
|  | 3 | Failed Tap | 22.69 | 34.31 | 31.185 | 3.6992 | 5002.0759 | 0.14391 | 1.3929 | 1.278 | 18.9659
|  | 4 | Building Toilet | 25.665 | NA | 34.045 | 2.8041 | 636.9231 | #VALUE! | #VALUE! | 0.42184 | 2.6414
|  | 3 | Failed Tap | 22.69 | 30.81 | 31.185 | 3.6992 | 5002.0759 | 1.2464 | 17.6346 | 1.278 | 18.9659
|  |  |  | Detected Value |  |  | Quantified Value
|  |  | Sample | Ct value All Legionella | Ct value Sg 1-15 | Ct value Sg 1 | Concentration HEX (cop/ul) | Concentration FAM (cop/ul) | Concentration TxRed (cop/ul) |  | Total copies of Legionella 1-15 in 1mL | Total copies of Legionella 2-15 in 1mL | Remarks based off acceptance criteria
|  | 1 | Water Tank | HEX | FAM | Texas Red | HEX | FAM | Texas Red
|  | 2 | 3plex-HW1 | NA | 39.11 | NA |  | 0.042852 |  |  | 0.021426 | 0.021426
|  | 3 | 3plex-HW2 | NA | 29.74 | NA |  | 38.318 |  |  | 19.159 | 19.159
|  | 4 | 3plex-HW3 | NA | 34.31 | NA |  | 1.3929 |  |  | 0.69644 | 0.69644
|  |  | 3plex-HW4 | NA | NA | NA |  | #VALUE! |  |  | #VALUE! | #VALUE!
|  |  | 3plex-WB | NA | 30.81 | NA |  | 17.6346 |  |  | 8.8173 | 8.8173
FAM
|  |  |  | 3plex-HW1 | 36.07 | 39.11
|  |  |  | 3plex-HW1 | 42.15
|  |  |  | 3plex-HW2 | 29.69 | 29.735
|  |  |  | 3plex-HW2 | 29.78
|  |  |  | 3plex-HW3 | 34.27 | 34.31
|  |  |  | 3plex-HW3 | 34.35
|  |  |  | 3plex-HW4 | N/A | #DIV/0!
|  |  |  | 3plex-HW4 | N/A
|  |  |  | 3plex-NC | N/A | #DIV/0!
|  |  |  | 3plex-NC | N/A
|  |  |  | 3plex-WB | 31.01 | 30.81
|  |  |  | 3plex-WB | 30.61

## I-Leg Count
Internal-Legionella Count | [10 empty] | linear regr.: | k | -3.3238 |  |  |  | Check Out These Column
[5 empty] | Eurofin Data | [5 empty] | HEX | d | 34.9852
[5 empty] | APHA 9222 B & H | ISO 11731:2017 | [5 empty] | y = kx + d
[5 empty] | EZ431 | EZ0EG Test | qPCR Data |  |  |  | Total count calculation (copy per uL) based on standards | [5 empty] | Final Concentration (Copy/mL )
|  | ID Discription | Remark | Sample ID | Escherichia coli Count | Total Legionella | Result (Legionella species) | Result (Legionella pneumophila 1-15) | Result (Legionella pneumophila 1) |  | (Legionella species) | (Legionella pneumophila 1-15) | (Legionella pneumophila 1) | (Legionella species) | (Legionella pneumophila 2-15) | (Legionella pneumophila 1) | Total copies of Legionella genomicDNA per mL (copy/mL)
[5 empty] | CFU/100mL | CFU/mL | Hex ct value | FAM ct value | TexasRed (ct value) | Remark | Log10(conc.) HEX | Log10(conc.) FAM | Log10(conc.) TxRed | Concentration HEX (cop/ul) | Concentration FAM (cop/ul) | Concentration TxRed (cop/ul) | (Legionella species) | (Legionella pneumophila 1-15) | (Legionella pneumophila 1)
| 04Oct2022; 20221005R1 | L8 TapWater (Female) |  | GS 001 | NA | 16 | 28.44 | 30.48 | 30.665 |  | 1.969 | 1.3553 | 1.2996 | 93.1136 | 22.6618 | 19.936 | 46.5568 | 11.3309 | 9.968
|  | L7 Shower (Female) |  | GIS 002 | NA | 8 | 26.65 | 28.66 | 29.1 |  | 2.5075 | 1.9028 | 1.7705 | 321.7519 | 79.9517 | 58.9464 | 160.8759 | 39.9759 | 29.4732
|  | Matrix |  |  | NA | <0.4 | 29.02 | 30.03 | N/A |  | 1.7945 | 1.4907 | #VALUE! | 62.3053 | 30.9509 | #VALUE! | 31.1526 | 15.4755 | #VALUE!
|  | Koi Pond |  |  | NA | 12 | 24.96 | 32.46 | 36.52 |  | 3.0159 | 0.75963 | -0.46179 | 1037.3962 | 5.7495 | 0.34531 | 518.6981 | 2.8747 | 0.17265
|  | NC (MiliQ) |  |  | <1 | <0.4 | N/A | N/A | N/A |  | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
|  | Water dispenser L7 (new) |  |  | <1 | <0.4 | 36.98 | N/A | N/A |  | -0.60018 | #VALUE! | #VALUE! | 0.25108 | #VALUE! | #VALUE! | 0.12554 | #VALUE! | #VALUE!
|  | Spiked in E.coli in NC(MiliQ) |  |  | <1 | 0.4 | 39.32 | N/A | N/A |  | -1.3042 | #VALUE! | #VALUE! | 0.049642 | #VALUE! | #VALUE! | 0.024821 | #VALUE! | #VALUE!
|  | EQUIPMENT STORE |  |  |  |  | 33.04 | 36.82 | N/A | 300mL | 0.58514 | -0.55205 | #VALUE! | 3.8471 | 0.28051 | #VALUE! | 1.9236 | 0.14026 | #VALUE!
|  | PROD WASH AREA-1 |  |  |  |  | 27.595 | 32.195 | 32.925 | 300mL | 2.2232 | 0.83935 | 0.61974 | 167.1957 | 6.908 | 4.1662 | 83.5978 | 3.454 | 2.0831
|  | PROD WASH AREA-2 |  |  |  |  | 29.29 | 33.58 | 35.53 | 300mL | 1.7133 | 0.42268 | -0.16396 | 51.677 | 2.6466 | 0.68555 | 25.8385 | 1.3233 | 0.34278
|  | PROD WASH AREA-3 |  |  |  |  | 31.555 | 37.23 | 36.08 | 300mL | 1.0319 | -0.67539 | -0.32942 | 10.7619 | 0.21116 | 0.46836 | 5.381 | 0.10558 | 0.23418
|  | MAIN KITCHEN COUNTER-1 |  | GIS |  |  | 28.93 | 34.385 | 36.38 | 300mL | 1.8216 | 0.18051 | -0.41968 | 66.3133 | 1.5153 | 0.38047 | 33.1566 | 0.75766 | 0.19024
|  | MAIN KITCHEN COUNTER-2 |  |  |  |  | 29.405 | 33.365 | 34.03 | 300mL | 1.6787 | 0.48736 | 0.2873 | 47.72 | 3.0716 | 1.9378 | 23.86 | 1.5358 | 0.96889
|  | MAIN KITCHEN STOVE-1 |  |  |  |  | 33.73 | 35.6 | N/A | 300mL | 0.37756 | -0.18502 | #VALUE! | 2.3854 | 0.6531 | #VALUE! | 1.1927 | 0.32655 | #VALUE!
|  | MAIN KITCHEN STOVE-2 |  |  |  |  | 30.465 | 34.625 | 36.125 | 300mL | 1.3598 | 0.1083 | -0.34296 | 22.8985 | 1.2832 | 0.45398 | 11.4493 | 0.64161 | 0.22699
|  | PASTRY AREA-1 |  |  |  |  | 30.935 | 33.41 | 35.625 | 300mL | 1.2184 | 0.47383 | -0.19254 | 16.5353 | 2.9773 | 0.64189 | 8.2676 | 1.4887 | 0.32095
|  | PASTRY AREA-2 |  |  |  |  | 30.175 | 33.3 | 34.7 | 300mL | 1.4471 | 0.50692 | 0.08574 | 27.9931 | 3.2131 | 1.2183 | 13.9966 | 1.6065 | 0.60913
|  | MAIN KITCHEN NEAR CHILLER-1 |  | Yan |  |  | 29.185 | 33.42 | 34.35 | 300mL | 1.7449 | 0.47082 | 0.19103 | 55.5758 | 2.9568 | 1.5525 | 27.7879 | 1.4784 | 0.77626
|  | MAIN KITCHEN NEAR CHILLER-2 |  |  |  |  | 30.955 | 33.895 | 34.32 | 300mL | 1.2124 | 0.32792 | 0.20006 | 16.3078 | 2.1277 | 1.5851 | 8.1539 | 1.0639 | 0.79256
|  | MAIN KITCHEN POINT |  |  |  |  | 28.98 | 33.975 | 34.4 | 300mL | 1.8066 | 0.30385 | 0.17599 | 64.0558 | 2.013 | 1.4997 | 32.0279 | 1.0065 | 0.74983
|  | MICRO KITCHEN SINK |  |  |  |  | 29.2 | 31.15 | 31.74 | 300mL | 1.7404 | 1.1537 | 0.97623 | 55.0013 | 14.2472 | 9.4675 | 27.5007 | 7.1236 | 4.7337
|  | MOTHER CARE ROOM |  |  |  |  | 28.38 | 30.15 | 30.605 | 300mL | 1.9871 | 1.4546 | 1.3177 | 97.0653 | 28.4822 | 20.7821 | 48.5326 | 14.2411 | 10.3911
|  | BEVERAGE COUNTER (ICE DISPENSER) |  |  |  |  | 26.37 | 30.545 | 31.575 | 300mL | 2.5918 | 1.3357 | 1.0259 | 390.6222 | 21.6641 | 10.6138 | 195.3111 | 10.832 | 5.3069
|  | SHOWER MALE |  |  |  |  | 30.23 | 34.41 | 38.17 | 300mL | 1.4305 | 0.17298 | -0.95818 | 26.9467 | 1.4893 | 0.11011 | 13.4733 | 0.74465 | 0.055054
[10 empty] | 300mL | 10.525 | 10.525 | 10.525 | 33494223639.8702 | 33494223639.8702 | 33494223639.8702 | 16747111819.9351 | 16747111819.9351 | 16747111819.9351
[10 empty] | TR(1/2) | 10.525 | 10.525 | 10.525 | 33494223639.8702 | 33494223639.8702 | 33494223639.8702 | 16747111819.9351 | 16747111819.9351 | 16747111819.9351
[10 empty] | FTR(1/2) | 10.525 | 10.525 | 10.525 | 33494223639.8702 | 33494223639.8702 | 33494223639.8702 | 16747111819.9351 | 16747111819.9351 | 16747111819.9351
[10 empty] | 300mL | 10.525 | 10.525 | 10.525 | 33494223639.8702 | 33494223639.8702 | 33494223639.8702 | 16747111819.9351 | 16747111819.9351 | 16747111819.9351
[10 empty] | 300mL | 10.525 | 10.525 | 10.525 | 33494223639.8702 | 33494223639.8702 | 33494223639.8702 | 16747111819.9351 | 16747111819.9351 | 16747111819.9351
[10 empty] | 300mL | 10.525 | 10.525 | 10.525 | 33494223639.8702 | 33494223639.8702 | 33494223639.8702 | 16747111819.9351 | 16747111819.9351 | 16747111819.9351
[10 empty] | (1/2) | 10.525 | 10.525 | 10.525 | 33494223639.8702 | 33494223639.8702 | 33494223639.8702 | 16747111819.9351 | 16747111819.9351 | 16747111819.9351
[10 empty] | 300mL | 10.525 | 10.525 | 10.525 | 33494223639.8702 | 33494223639.8702 | 33494223639.8702 | 16747111819.9351 | 16747111819.9351 | 16747111819.9351
[10 empty] | 300mL | 10.525 | 10.525 | 10.525 | 33494223639.8702 | 33494223639.8702 | 33494223639.8702 | 16747111819.9351 | 16747111819.9351 | 16747111819.9351
[10 empty] | 300mL | 10.525 | 10.525 | 10.525 | 33494223639.8702 | 33494223639.8702 | 33494223639.8702 | 16747111819.9351 | 16747111819.9351 | 16747111819.9351
FALSE

## R-Leg Count
Reduced Legionella Count | [10 empty] | linear regr.: | k | -3.3238 |  |  |  | Check Out These Column
[5 empty] | Eurofin Data | [5 empty] | HEX | d | 34.9852
[5 empty] | APHA 9222 B & H | ISO 11731:2017 | [5 empty] | y = kx + d
[5 empty] | EZ431 | EZ0EG Test | qPCR Data |  |  |  | Total count calculation (copy per uL) based on standards | [5 empty] | Final Concentration (Copy/mL )
|  | ID Discription | Remark | Sample ID | Escherichia coli Count | Total Legionella | Result (Legionella species) | Result (Legionella pneumophila 2-15) | Result (Legionella pneumophila 1) |  | (Legionella species) | (Legionella pneumophila 2-15) | (Legionella pneumophila 1) | (Legionella species) | (Legionella pneumophila 2-15) | (Legionella pneumophila 1) | Total copies of Legionella genomicDNA per mL (copy/mL)
[5 empty] | CFU/100mL | CFU/mL | Hex ct value | FAM ct value | TexasRed (ct value) | Remark | Log10(conc.) HEX | Log10(conc.) FAM | Log10(conc.) TxRed | Concentration HEX (cop/ul) | Concentration FAM (cop/ul) | Concentration TxRed (cop/ul) | (Legionella species) | (Legionella pneumophila 2-15) | (Legionella pneumophila 1)
|  | L7 Shower (Female) |  | GIS 002 | NA | 8 | 26.65 | 28.66 | 29.1 |  | 2.5075 | 1.9028 | 1.7705 | 321.7519 | 79.9517 | 58.9464 | 160.8759 | 39.9759 | 29.4732
| 04Oct2022; 20221005R1 | L8 TapWater (Female) |  | GS 001 | NA | 16 | 28.44 | 30.48 | 30.665 |  | 1.969 | 1.3553 | 1.2996 | 93.1136 | 22.6618 | 19.936 | 46.5568 | 11.3309 | 9.968
|  | level 7 (E lysis) ctrl |  |  |  | 8 (Ref) | 28.67 | N/A | N/A | 300mL | 1.8998 | #VALUE! | #VALUE! | 79.3998 | #VALUE! | #VALUE! | 39.6999 | #VALUE! | #VALUE!
|  | level 7 (E lysis) PMA |  |  |  |  | 29.12 | N/A | N/A | 300mL | 1.7644 | #VALUE! | #VALUE! | 58.1354 | #VALUE! | #VALUE! | 29.0677 | #VALUE! | #VALUE!
|  | level 7 (E lysis) PMA + Enhancer |  |  |  |  | 29.24 | N/A | N/A | 300mL | 1.7283 | #VALUE! | #VALUE! | 53.4982 | #VALUE! | #VALUE! | 26.7491 | #VALUE! | #VALUE!
|  | level 8 (E lysis) ctrl |  | GIS |  | 16 (Ref) | 29.48 | N/A | N/A | 300mL | 1.6561 | #VALUE! | #VALUE! | 45.3041 | #VALUE! | #VALUE! | 22.652 | #VALUE! | #VALUE!
|  | level 8 (E lysis) PMA |  |  |  |  | 29.85 | N/A | N/A | 300mL | 1.5448 | #VALUE! | #VALUE! | 35.0611 | #VALUE! | #VALUE! | 17.5305 | #VALUE! | #VALUE!
|  | level 8 (E lysis) PMA + Enhancer |  |  |  |  | 30.48 | N/A | N/A | 300mL | 1.3553 | #VALUE! | #VALUE! | 22.6618 | #VALUE! | #VALUE! | 11.3309 | #VALUE! | #VALUE!
|  | level 8 (LB) ctrl |  |  |  |  | 29.78 | N/A | N/A | 300mL | 1.5659 | #VALUE! | #VALUE! | 36.8031 | #VALUE! | #VALUE! | 18.4016 | #VALUE! | #VALUE!
|  | level 8 (LB) PMA |  |  |  |  | 32.18 | N/A | N/A | 300mL | 0.84386 | #VALUE! | #VALUE! | 6.9801 | #VALUE! | #VALUE! | 3.4901 | #VALUE! | #VALUE!
|  | level 8 (LB) PMA + Enhancer |  |  |  |  | 34.74 | N/A | N/A | 300mL | 0.073706 | #VALUE! | #VALUE! | 1.185 | #VALUE! | #VALUE! | 0.59248 | #VALUE! | #VALUE!
| 2011-11-01 | Level 7 (E lysis) Ctrl |  |  |  | 8 (Ref) | 32.31 | N/A | N/A | 300mL | 0.80475 | #VALUE! | #VALUE! | 6.379 | #VALUE! | #VALUE! | 3.1895 | #VALUE! | #VALUE!
|  | Level 7 (E lysis) PMA |  |  |  |  | 31.25 | N/A | N/A | 300mL | 1.1236 | #VALUE! | #VALUE! | 13.2937 | #VALUE! | #VALUE! | 6.6469 | #VALUE! | #VALUE!
|  | Level 7 (E lysis) PMA+Enhancer |  |  |  |  | 30.02 | N/A | N/A | 300mL | 1.4937 | #VALUE! | #VALUE! | 31.1661 | #VALUE! | #VALUE! | 15.583 | #VALUE! | #VALUE!
|  | Level 7 Dead (E lysis) Ctrl |  |  |  |  | 32.84 | N/A | N/A | 300mL | 0.64531 | #VALUE! | #VALUE! | 4.4188 | #VALUE! | #VALUE! | 2.2094 | #VALUE! | #VALUE!
|  | Level 7 Dead (E lysis) PMA |  |  |  |  | 37.37 | N/A | N/A | 300mL | -0.71751 | #VALUE! | #VALUE! | 0.19164 | #VALUE! | #VALUE! | 0.095821 | #VALUE! | #VALUE!
|  | Level 7 Dead (E lysis) PMA+Enhancer |  |  |  |  | 35.9 | N/A | N/A | 300mL | -0.27527 | #VALUE! | #VALUE! | 0.53055 | #VALUE! | #VALUE! | 0.26528 | #VALUE! | #VALUE!
|  | Level 7 (LB) Ctrl |  |  |  |  | 30.085 | N/A | N/A | 300mL | 1.4741 | #VALUE! | #VALUE! | 29.7939 | #VALUE! | #VALUE! | 14.897 | #VALUE! | #VALUE!
|  | Level 7 (LB) PMA |  |  |  |  | 30.85 | N/A | N/A | 300mL | 1.244 | #VALUE! | #VALUE! | 17.5381 | #VALUE! | #VALUE! | 8.7691 | #VALUE! | #VALUE!
|  | Level 7 (LB) PMA+Enhancer |  |  |  |  | 34.11 | N/A | N/A | 300mL | 0.26324 | #VALUE! | #VALUE! | 1.8333 | #VALUE! | #VALUE! | 0.91666 | #VALUE! | #VALUE!
|  | Level 7 Dead (LB) Ctrl |  |  |  |  | 29.44 | N/A | N/A | 300mL | 1.6682 | #VALUE! | #VALUE! | 46.5769 | #VALUE! | #VALUE! | 23.2885 | #VALUE! | #VALUE!
|  | Level 7 Dead (LB) PMA |  |  |  |  | 40.78 | N/A | N/A | 300mL | -1.7434 | #VALUE! | #VALUE! | 0.018056 | #VALUE! | #VALUE! | 0.0090279 | #VALUE! | #VALUE!
|  | Level 7 Dead (LB) PMA+Enhancer |  |  |  |  | N/A | N/A | N/A | 300mL | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
FALSE

## Ec 078 Stds
FAM
UidA (All Ecoli)
Extracted E.coli (078) | qPCR | dPCR |  |  | Standard FAM Ct value |  |  | qPCR
Standards | Mean (ct value) | Concentration ( Copies/ul) | Log10 Copies/ul |  | R1 | R2 | R3 | Mean (ct value) | Concentration ( Copies/ul) | Mean value (ct)
S1 | 23.1 | 2475.8 | 3.3937 |  | 23.13 | 23.15 | 23.02 | 23.1 | 511 | 24.415
S2 | 26.3133 | 247.8 | 2.3941 |  | 26.17 | 26.54 | 26.23 | 26.3133 | 50.57 | 28.0833
S3 | 29.6667 | 23.4 | 1.3692 |  | 29.74 | 29.81 | 29.45 | 29.6667 | 5 | 31.57
S4 | 33.9333 | 2.698 | 0.43104 |  | 33.63 | 33.97 | 34.2 | 33.9333 | 0.44 | 34.9133
|  |  | Amplification Efficiency = 10^(-1/slope) - 1 |  |  | 89.19%
|  |  |  | linear regr.: | k | -3.6115
[5 empty] | d | 35.1044
y = kx + d

## Ec Count
Ecoli Count | [7 empty] | linear regr.: | k |  | -3.6115 |  | Check Out These Column
|  |  |  | Eurofin Data |  |  |  | HEX | d |  | 35.1044
|  |  |  | APHA 9222 B & H | ISO 11731:2017 |  |  |  | y = kx + d
|  |  |  | EZ431 | EZ0EG Test | qPCR Data |  | Total count calculation (copy per uL) based on standards |  |  |  |  | Final Concentration (Copy/mL )
| Sample ID Discription | Remark | Sample ID | Escherichia coli Count | Total Legionella | E.coli Assay |  | (DH5a) | (DH5a) | 150uL for elution | Total copies of Ec genomicDNA per mL (copy/mL)
|  |  |  | CFU/100mL | CFU/mL | Famct value | Remark | Log10(conc.) Fam | Concentration Fam (cop/ul) | Copy in sample (By Calculation) | (Ec species)
| Standard (Ec 078 1000x) |  |  |  |  | 23.1 |  | 3.3239 | 2108.3368 | 316250.527 | 1054.1684
| Standard (Ec 078 10k xd) |  |  |  |  | 26.3133 |  | 2.4342 | 271.7633 | 40764.5002 | 135.8817
| Standard (Ec 078 100k xd) |  |  |  |  | 29.6667 |  | 1.5057 | 32.0388 | 4805.8253 | 16.0194
| Standard (Ec 078 1M xd) |  |  |  |  | 33.9333 |  | 0.32427 | 2.1099 | 316.4888 | 1.055
20221122 | Live Elysis Control |  |  |  |  | 23.425 |  | 3.2339 | 1713.7592 | 257063.8871 | 856.8796 | 51412.7774
20221122 | Live Elysis PMA |  |  |  |  | 25.14 |  | 2.7591 | 574.2203 | 86133.0395 | 287.1101 | 17226.6079
20221122 | Live Elysis PMA + Enhancer |  |  |  |  | 26.63 |  | 2.3465 | 222.0793 | 33311.8977 | 111.0397 | 6662.3795
20221122 | Live LB control |  |  |  |  | 24.12 |  | 3.0415 | 1100.2953 | 165044.2944 | 550.1476 | 33008.8589
20221122 | Live LB PMA |  |  |  |  | 27.585 |  | 2.0821 | 120.8024 | 18120.3537 | 60.4012 | 3624.0707
20221122 | Live LBPMA + Enhancer |  |  |  |  | 29.44 |  | 1.5684 | 37.0202 | 5553.0345 | 18.5101 | 1110.6069
#REF!

## Sample list
| Colletection Date | Sample pick-up Date | Sample submission form | ID Discription | Water Features | Sample ID | Legionella Test | E.Coli Test | Quotation number
1 | 2022-10-04 | 2022-10-04 |  | L8 TapWater (Female) | Ccooling tower | GS 001 | Yes | No | US9PAP220249-01
2 | 2022-10-04 | 2022-10-04 |  | L7 Shower (Female) | Cooling tower | GS 002 | Yes | No
3 | 2022-10-04 | 2022-10-04 |  | Matrix | Cooling tower | GS 003 | Yes | No
4 | 2022-10-04 | 2022-10-04 |  | Koi Pond | Pond | GS 004 | Yes | No
5 | 2022-10-04 | 2022-10-04 |  | NC (MiliQ) | Drinking water | GS 005 | Yes | Yes
6 | 2022-10-04 | 2022-10-04 |  | Water dispenser L7 (new) | Drinking water | GS 006 | Yes | Yes
7 | 2022-10-04 | 2022-10-04 |  | Spiked in E.coli in NC(MiliQ) | Drinking water | GS 007 | Yes | Yes
[6 empty] | 22 | L | 1
[6 empty] | 22 | E | 1

## 19Oct
| Pre cleaning |  |  | Pre cleaning
| Eurofin Data | GIS Data |  | Post cleaning | GIS Data
Sample Description | Result (Legionella | Result (Legionella | Result (Legionella | Result (Legionella | Result (Legionella | Result (Legionella
| pneumophila 1) | pneumophila 1) | pneumophila 2-15) | pneumophila 1) | pneumophila 1) | pneumophila 2-15)
LEVEL 30E MAIN KITCHEN - EQUIPMENT STORE | 3 |  |  | 3 | NA | NA
LEVEL 30E MAIN KITCHEN - PROD WASH AREA-1 | 2 |  |  | 3 | 2.86 | 0.37
LEVEL 30E MAIN KITCHEN - PROD WASH AREA-2 | 3 |  |  | 1 | 0.47 | 0.71
LEVEL 30E MAIN KITCHEN - PROD WASH AREA-3 | 3 |  |  | <1 | 0.32 | NA
LEVEL 30E MAIN KITCHEN - MAIN KITCHEN COUNTER-1 | 2 |  |  | <1 | 0.26 | 0.4
LEVEL 30E MAIN KITCHEN - MAIN KITCHEN COUNTER-2 | 2 |  |  | <1 | 1.33 | 0.05
LEVEL 30E MAIN KITCHEN - MAIN KITCHEN STOVE-1 | 2 |  |  | <1 | NA | NA
LEVEL 30E MAIN KITCHEN - MAIN KITCHEN STOVE-2 | <0.01 |  |  | <1 | 0.31 | 0.24
LEVEL 30E MAIN KITCHEN - PASTRY AREA-1 | 2 |  |  | <1 | 0.44 | 0.89
LEVEL 30E MAIN KITCHEN - PASTRY AREA-2 | 2 |  |  | <1 | 0.84 | 0.61
LEVEL 30E MAIN KITCHEN - MAIN KITCHEN NEAR CHILLER-1 | 3 |  |  | 1 | 1.07 | 0.26
LEVEL 30E MAIN KITCHEN - MAIN KITCHEN NEAR CHILLER-2 | 2 |  |  | 2 | 1.09 | NA
LEVEL 30E MAIN KITCHEN - MAIN KITCHEN POINT | 2 |  |  | <1 | 1.03 | NA
LEVEL 30E MAIN KITCHEN - DISH WASH AREA-1 | 2 |  |  | 4 | - | -
LEVEL 30E MICRO KITCHEN - MICRO KITCHEN SINK | 3 |  |  | 5 | 6.47 | 0.42
LEVEL 30E MOTHERS ROOM - MOTHERS ROOM | 2 |  |  | <0.01 | 14.14 | 0.09
LEVEL 30E BEVERAGE COUNTER - BEVERAGE COUNTER (ICE DISPENSER) | <0.01 |  |  | - | 7.25 | 3.44
LEVEL 30E TOILET - SHOWER ROOM (MALE TOILET) | <0.01 |  |  | <0.01 | 0.08 | 0.57
LEVEL 18W MICRO KITCHEN - MICRO KITCHEN | <0.01 |  |  | <0.01 | - | -
LEVEL 29E MOTHERS ROOM - MOTHERS ROOM | <0.01 |  |  | - | - | -
Water tank (Unknown) | - | 10.19 | 4.72
Passed Tap |  | 2.93 | 4.87
Failed Tap |  | 9.48 | 242.04
Building Toilet |  | 1.32 | 2.66